## Slide 1: Mzkit User Manual
BioNovoGene Bio-Pharm Technology Co., Ltd
Present: BioDeep R&D Department
E-Mail: R&D@biodeep.cloud
BioNovoGene m/z Data Toolkit User Manual
January, 2021, SuZhou, China

## Slide 2: Introduction
The main function of Mzkit includes:
Raw Data Viewer
Metabolite Feature Annotation
Molecular Networking
MS-Imaging
Targeted Quantification
Open Source M/z Raw Data File Format Supports:
mzML, mzXML, imzML, mzPack, netCDF, HDF, Mgf

## Slide 3: LICENSE
Mzkit application and its source code is distributed under the MIT open source license
Source Code:
https://github.com/xieguigang/mzkit
Downloads:
http://www.biodeep.cn/downloads?lang=en-US
This document is distributed under Creative Commons Attribution-ShareAlike 4.0 International Public License

## Slide 4: Screenshots
c. 2D Scatter
d. MS-Imaging
a. Chromatography
b. Spectrum

## Slide 5: Main Screen
The main screen page of Mzkit application.
Data Explorer
Content Viewer
Content Properties
Status Toolstrip
Toolkit Ribbon

## Slide 6: Raw Data Decode

## Slide 7: File Formats
mzXML
mzML
imzML
CDF
HDF
mzPack
Raw
Wiff
…
Open Source Format
Various Closed Source
Vendor Format

## Slide 8: mzXML demo
MS scan meta data
Steps of decode MS scan raw data
base64 decode
decompress based on the compressionType option
BitConvert decode to float or double
create matrix

## Slide 9: Raw Decode
Demo code of Base64Decoder for decode mass spectrum
content data in VisualBasic language.
Demo code of decode mass spectrum content data in R#
programming language by using the “buffer” package
module.
base64 decode
zlib/gzip decompression
convert to network byteorder
bit convert to floats

## Slide 10: View Untargeted Data
TIC, BPC, EIC
Raw Scatter 2D
XIC Peaks / Contour Plot
PDA
UV Scan

## Slide 11: Introduction
Mzkit supports for view data content in common used open source mzXML and mzML file.
Although the raw data file is in various format, but the mass spectrum data in both of them can be unify into the identical data model for downstream visualization or data analysis in Mzkit.

## Slide 12
Imports Raw Data
Due to the reason of the data size of untargeted metabolomics raw data file is very large, and reading the raw data in text file is usually slowly. So you should always imports the raw data file into Mzkit at first and then you can view the raw data content in Mzkit.
Open Single File
The file open button accepts many file formats, but only supports one file in one click
Imports Multiple Files
Use the context menu on the file explorer for imports a set of raw data files in batch
A progress bar dialog will appear if this is the first file
that imports into mzkit
The imports progress will be running in background is the first file have been
imported into mzkit. You can view the progress in task list tab.

## Slide 13: View 2D Scatter Snapshot
The 2D scatter snapshot of a raw data file is consist with m/z and rt these two dimension data plot. And the scatter color is mapping to the ms1 ion intensity. Just click on the file node on the file explorer then you will see the 2D scatter snapshot of your raw data file.
Select File Node
Just click on the file node to view its 2D scatter snapshot

## Slide 14: View Contour Snapshot
The contour snapshot of a raw data file is consist with m/z and rt these two dimension data as 2D point. And the contour level is mapping to the ms1 ion intensity. Just click on the file node on the file explorer then you will see the contour scatter snapshot of your raw data file.
Select File Node
Just click on the file node, then select the contour plot on context menu to view its 2D contour snapshot

## Slide 15: View 3D XIC Snapshot
The 3D XIC snapshot of a raw data file is consist with m/z and rt these two dimension data plot. And the intensity value is mapping to the height of the ions XIC peak. Just click on the file node on the file explorer then you will see the 3D XIC snapshot of your raw data file.
Select File Node
Just click on the file node, then select the XIC peaks on context menu to view its 2D scatter snapshot

## Slide 16: View TIC/BPC
Both the TIC and BPC plot is a kind of time-signal plot but with different meanings.
TIC is the chromatography plot of total ions.
BPC is the Chromatography plot of the max intensity ion.
View Chromatography
View TIC/BPC from the contex menu of the file explorer. If multiple file was selected, then TIC/BPC overlaps will be display

## Slide 17: View Scan Feature
The raw data is consists with multiple spectrum scans in MS1 or MS2 levels
The scan in a raw data file is organized in a tree liked data: A parent MS1 scan contains multiple product ion scans.
View Mass Spectrum
Click one of the scan node in the Feature explorer for view its mass spectrum data.

## Slide 18: View XIC
XIC plot is a kind of time-signal plot which is very similar to TIC or BPC.
XIC comes from a set of specific product ion with their parent intensity correlation to scan times.
View XIC Chromatography
Click one of the Ms2 scan node in the Feature explorer, and then select Show XIC on context menu for view its Chromatography plot

## Slide 19: View PDA
PDA plot is a kind of time-signal plot which is very similar to the TIC plot.
Electromagnetic radiation spectrum is supported in mzML file, mzXML file format is not supprted.
Show PDA
Just select the Show PDA menu item on the context menu in the UV scans explorer

## Slide 20: View UV scans
Electromagnetic radiation spectrum is supported in mzML file, mzXML file format is not supprted.
Show UV Overlap
Just select the Show UV Overlap menu item on the context menu in the UV scans explorer. If multiple scan was selected, then UV scan overlap will be plot

## Slide 21: DIA Annotation

## Slide 22: Formula Search
Formula Link
Click on the link of one formula for query in BioDeep knowledge database
Formula Search
Select one of MS2 ion feature for run formula search
View Search Result
Formula search result includes formula, exact mass, ppm error between the m/z and the evaluate m/z, precursor info

## Slide 23: DEMO Search
One Ms2 ion feature with m/z 132.1016 and rt at 112 sec was selected for run formula de-novo search.
A list of candidate formula was generated, with different exact mass and precursor type for match the parent ion m/z 132.1016

## Slide 24
Query BioDeep
Click of the formula link for query BioDeep KB with formula as search term
Find Leucine match the formula search result:
C6H13NO2[M+H]+

## Slide 25: DIA Annotation (MetaDNA)

## Slide 26
MetaDNA Search
MetaDNA Result Table
Show the MetaDNA DIA annotation result at here, include KEGG ID, KEGG reaction chain information, KEGG partner compound information
MetaDNA Search
Run metaDNA search via context menu on the raw data file features explorer

## Slide 27
View MetaDNA Infer Alignment
MetaDNA Alignment
The MS2 DIA alignment from metaDNA result will be show at the plot viewer
The L-Methionine its predicted as 104.18/133.12 in MS2
View MetaDNA Alignment
Select one row in the result table, and then show contex menu with mouse right button click

## Slide 28
BioDeep MS Reference(MoNA)
One of the MS reference data record(MoNA: CE000446) in BioDeep database shows that the MS2 of L-Methionine is [104.0528, 133.03]
The metaDNA DIA annotation result is [104.18, 133.12], the metaDNA DIA annotation result is agree with the MS reference data record.
ROI feature M150T105 is annotated as L-Methionine successfully via metaDNA DIA annotation in mzkit software.

## Slide 29: Chromatography Explorer

## Slide 30: Chromatography Explorer
Open Explorer
Select one of the raw data file in the file explorer and then open viewer from the context menu.
Select RT range
View all MS1 ions in the given range of MS1 scans by rtmin and rtmax.
View MS1
Display MS1 ions in a given rt range.

## Slide 31: Explorer Contex Menu
Reset: for reset explorer to initial status
Pin: mouse select of the RT range will not update explorer to new RT range
TIC/BPC: switch the Chromatography between the TIC and BPC
Filter MS2: Show the ms2 ions in the given RT selection range

## Slide 32: View MS1 Ions
Show Matrix
View all MS1 ions in the given range of MS1 scans by rtmin and rtmax.

## Slide 33: MS-Alignment

## Slide 34: Cos Similarity
vector x, y could be in different value scale range.
If the two vector in the same direction, then it means the angle of this two vector is ZERO, projection between the vector will be maximized. cos(0°)=1, means of those two spectra is identical;
If the two vector is not relevant to each other, which means orthogonal, then the angle of this two vector should be 90, projection between the vector is ZERO. cos(90°)=0, means the two spectra are totally different.

## Slide 35
MS-Alignment
m/z | intensity
100 | 100
200 | 50
300 | 80
350 | 80
m/z | intensity
100 | 100
300 | 60
310 | 10
reference
Forward: using query as base matrix
m/z | 100 | 200 | 300 | 350
query | 100 | 50 | 80 | 80
reference | 100 | 0 | 60 | 0
Reverse: using reference as base matrix
m/z | 100 | 300 | 310
query | 100 | 80 | 0
reference | 100 | 60 | 10
Cos similarity result i, this two spectra is not so similar
query

## Slide 36
MS-Alignment
Cos similarity result i, this two spectra is not so similar

## Slide 37
MS2 Search
Spectrum Search
Click Spectrum Search Icon for open MS2 search toolkit
Input Spectrum
Input MS2 spectrum matrix at the input table
Run Search
Click Search button for run MS2 search between all raw data files in the file explorer

## Slide 38
View MS2 Search Result
Search Result
All of the ms2 will be align with the custom MS matrix. No hits means no similar MS matrix in raw data file.
View MS alignment
Mouse click with right button will pop up a context menu for view MS2 alignment result

## Slide 39
View MS2 Alignment
MS2 Alignment
View the ms2 alignment result between the custom MS spectrum input and target ions in raw data file
Alignment Score
View the alignment score details

## Slide 40: MS-Imaging Viewer

## Slide 41: Introduce MS-Imaging
Mass spectrometry imaging (MSI) is a technique used in mass spectrometry to visualize the spatial distribution of molecules, as biomarkers, metabolites, peptides or proteins by their molecular masses. By choosing a peak in the resulting spectra that corresponds to the compound of interest, the MS data is used to map its distribution across the sample. This results in pictures of the spatially resolved distribution of a compound pixel by pixel.
Römpp A, Spengler B. Mass spectrometry imaging with high resolution in mass and space.
Histochemistry and Cell Biology. 2013 Jun;139(6):759-783.
DOI: 10.1007/s00418-013-1097-6.

## Slide 42: imzML for MS-imaging
Mzkit Supports the imzML file format for view ms-imaging content.
[Main] Open
Select imzML format
Select ion for view image
Steps for view MS-imaging
in Mzkit:

## Slide 43: MS-imaging
MS-imaging in Mzkit
Ion Explorer
Select one ion or multiple ions at here.
Context menu on mouse right click for run image rendering.
Render Options
Tweaks of the MS-Imaging rendering options
MSI Viewer
View MS-Imaging rendering image output

## Slide 44
MSI Ion Rendering (RGB)
MSI ions rendering in RGB channels composition mode only supports 3 ion layers due to the reason of color is compose by RGB channels.
Ion Explorer
Select 3 ion at here.
Context menu on mouse right click for run image rendering.

## Slide 45
MSI Multiple Ion Rendering
Another mode for single/multiple ions MSI rendering in mzkit has no number limitation of the ions.
Ion Explorer
Select any ions at here.
Context menu on mouse right click for run image rendering.

## Slide 46
View Ion MS Content
Click on any pixel of the MSI rendering image for display the MS ions on target pixel scan
MSI Viewer
Click on the image for view MS ions on target pixel scan or select a region on image for view the total MS ion content on that region.

## Slide 47: View GC-MS Targeted

## Slide 48: View GC-MS Targeted
The GC-MS targeted data in CDF file format is supported in Mzkit.
View GC-MS targeted data in Mzkit includes view TIC data and SIM ions for metabolite quantification.
[Main] Open
Select CDF format
View GC-MS data in Mzkit
Steps for view GC-MS targeted data
in Mzkit:

## Slide 49: View GC-MS Targeted
View SIM Ions Of GC-MS Targeted
Feature Explorer
Select one ion at here for view mass spectrum plot data.
Select TIC node for view TIC plot of the current GC-MS experiment
Peak Explorer
View Ion peak on TIC at here, and select rt range with mouse drag.
MS Viewer
View the mass spectrum of the selected ion peak

## Slide 50: View GC-MS XIC
View SIM Ions Of GC-MS Targeted
Feature Explorer
Select one ion at here for view its XIC plot.

## Slide 51: View LC-MS/MS Targeted

## Slide 52: View LC-MS/MS Targeted
MRM stands for Multiple Reaction Monitoring.
In MRM, a specific ion (precursor ion) is selected in Q1 (first stage MS) from among various ions ionized by the ionization probe, the ion is destructed (collision-induced dissociation) in the collision cell (Q2, collision chamber), and a specific ion is detected from among the destructed ions (product ions) in Q3 (second stage MS).
Multiple channels can be set in one measurement.
[Main] Open
Select mzML format
View LC-MS/MS data in Mzkit
Steps for view LC-MS/MS targeted data
in Mzkit:

## Slide 53: View LC-MS/MS Targeted
View LC-MS/MS Targeted MRM data in Mzkit
Feature Explorer
Select one ion pair at here for view extracted ion chromatography
plot data.
Select TIC node for view TIC plot of the current LC-MS/MS experiment
Content Viewer
View the mass spectrum of the selected ion peak

## Slide 54: MRM TIC/BPC Overlaps
View LC-MS/MS Targeted MRM data in Mzkit
Feature Explorer
Select multiple MRM raw data files for view TIC/BPC overlaps
TIC Overlaps
View the TIC/BPC overlaps at here
TIC/BPC
Show TIC/BPC overlaps via context menu

## Slide 55: Targeted quantification analysis

## Slide 56: Create Linear Models
Targeted linear analysis tools in mzkit supports LC-MSMS(MRM) and GC-MS(SIM) data analysis.
Steps for create linear model in mzkit:
imports reference linear raw data files
edit reference contents for each level
setup IS relationship
[Tools] Targeted Quantification
Imports Linear Reference
Edit content levels and IS relationships

## Slide 57: Create Linear Models
Linear Viewer
view linear modelling result
Linear Modeller
Edit linear content levels and IS relationship with the target metabolites

## Slide 58: 感谢观看
苏州帕诺米克生物医药科技有限公司
Thank you for your attention
地址：江苏省苏州市工业园区新平街388号22幢2层
电话：0512-62959105
邮箱：info@bionovogene.com
网址：https://www.bionovogene.com www.biodeep.cn
科研服务公众号
健康服务公众号
科研服务订阅号